## Feuil1
Error Type | Descpription | Possible Causes | Alarm |  | Error Handling | Alarm Response
Open Circuit/ Disconnected | Thermocouple input missing or disconnected | Broken wire, loose terminal, thermocouple unplugged |  | Sensor Fault Alarm | 1. Flag for maintenance 2. Use pre determined value as thermocouple input until replaced. | Show "OPEN"
Sensor Shorted | Thermocouple inputs shorted. E.g wire to wire, to grounded surfaces | Pinched wire, installation damage |  | Sensor Fault Alarm | 1. Flag for maintenance 2. Use pre determined value as thermocouple input until replaced. | Show "SHORT"
Out of range reading | Reading outside Thermocouple range (Type N: -40°C to 1300°C ) | Faulty thermocouple, signal distortion |  | Input Range Alarm | 1. Flag for maintenance | Show "Invalid Value"
2. Use pre determined value as thermocouple input until replaced
High Temperature Warning | Temperature exceeds defined warning threshold | Normal operating temperature approacing upper limits, |  | High Temperature Warning | 1.Flag for maintenace and phycisist during post processing of data. | Show Visual/ audible warning on HMI (highlight temperature/ thermocouple ID)
Overtemperature Fault | Critical temperature limit exceeded | Cooling failure, runaway process |  | Overtemperature Alarm |  | Raise alarm on HMI for operator action.
Communication Failure | Link to PLC,CPU,or CODAC is broken | Cable unplugged, CODAC network down |  | Communication Alarm | 1. Flag for maintenance 2. Use pre determined value as thermocouple input until replaced | "No Data" displayed on HMI
Drift or Noise Detected | Fluctuation in readings with constant temperature input | Ground loop, EMI |  | Sensor Stability Warning | 1. Confrim drift or wear using post processing data. 2. Flag for replacement if confirmed.
Calibration error | Values constantly offset from known standard | Thermocouple degraded/failing |  | Calibration Warning | 1. Confrim drift or wear using post processing data. 2. Flag for replacement if confirmed. | Recommend recalibration/ replacement
No Image received | A DAQ Fast Controller does not receive images during acquisition process | Camera failure, frame grabber failure, MTCA chassis failure, Camera power supply failure, DAQ channel communication failure |  | Acquisition Error | 1. Flag for Maintenance 2. Investigation to identify the faulty device(s) | Raise alarm on HMI for operator action.
IR Camera Failure | Camera dysfunctioning or not functioning anymore while other cameras are functioning | ageing, overvoltage… |  | IR camera Error | 1. Flag for Maintenance 2. If camera model redundancy and mission ongoing, switch on and configure the second camera to resume mission 3. otherwise, a. analyse error code if any b.faulty camera power cycle 4. Flag for replacement if confirmed | Raise alarm on HMI for operator action.
Visible Camera Failure | Camera dysfunctioning or not functioning anymore while other cameras are functioning | ageing, overvoltage… |  | VIS camera Error | 1. Flag for Maintenance 2.Analyse error code if any 3.Faulty camera power cycle 4. Flag for replacement if confirmed | Raise alarm on HMI for operator action.
Frame Grabber failure | No image received although Camera and Camera power supply functioning whereas images from other DAQ channels can be received | Frame grabber failure |  | Frame grabber Error | 1. Flag for Maintenance 2. Fault cause identification 3. Flag for replacement if confirmed | Raise alarm on HMI for operator action.
MTCA chassis failure | MTCA chassis fault | MTCA chassis failure |  | MTCA chassis Error | 1. Flag for Maintenance 2. Analyse Error codes if any 3. Fault cause identification 4. Flag for replacement if confirmed | Raise alarm on HMI for operator action.
Media Converter / Power Supply failure | No image received altough camera and Frame grabber well functioning | Media converter dysfunctioning, Camera power supply failure or broken cable/wire Power supply failure, IO board failure, failure of an used channel of the IO board |  | Image Transfer Error | 1. Flag for Maintenance 2. Analyse Error codes if any 3. Fault cause identification 4. Flag for replacement if confirmed | Raise alarm on HMI for operator action.
DAQ channel communication failure | No more communication between a MTCA chassis and its DAQ Fast Controller | NAT-MCH-PHYS80 board failure, MPCIE16 board failure, Broken cable / wire |  | DAQ channel communication Error | 1. Flag for Maintenance 2. Analyse Error codes if any 3. Fault cause identification 4. Flag for replacement if confirmed | Raise alarm on HMI for operator action.
Software execution error | The software fails or raises an exception | Division per zero, Array length… |  | Software Error | 1. Flag for Maintenance 2. Analyse potential error source functions | Raise error on HMI.
Acquisition parameters not available | Configuration data not available or incomplete | Configuration data loading failure |  | Acquisition Data Error | 1. Reset connection with PON 2. Re-load configuration 3. Re-apply configuration 4. If error persistent, flag for Maintenance | Raise error on HMI.
Faulty acquisition parameters | Acquisition parameters not allowed by the camera model | Faulty configuration data |  | Faulty acquisition parameters Error | 1. Flag for new configuration data 2. Wait for new configuration available 3. Load new configuration 4. Apply new configuration | Raise alarm on HMI for operator action.
Calibration parameters not available | Configuration data not available or incomplete | Configuration data loading failure |  | Calibration Data Error | 1. Flag for new configuration data 2. Wait for new configuration available 3. Load new configuration 4. Apply new configuration | Raise alarm on HMI for operator action.
Algorithm parameters not available | Configuration data not available or incomplete | Configuration data loading failure |  | Algorithm Data Error | 1. Reset connection with PON 2. Re-load configuration 3. Re-apply configuration 4. If error persistent, flag for Maintenance | Raise error on HMI.
IR camera overheat | The camera temperature is out of wellfunctioning range given by the manufacturer | Cooling System malfunctioning |  | Image Quality Degradation Warning | 1. Flag for Maintenance | Raise warning on HMI.
IR camera type not adapted to Measurement Parameter calculation | The system configuration plans a measurement parameter calculation to be performed by not adapted camera type (HR or HS) | Error in configuration |  | IR camera model not adapted Warning | Measurement Parameter calculation degraded | Raise warning on HMI
CPU load close to maximum | High CPU load | Important computational load |  | CPU load Warning | 1. Follow up CPU load | Raise warning on HMI
CPU overload | Maximum CPU reached. Risk of latency increase | Very important computational load |  | CPU overload Warning | 1. Follow up CPU load | Raise warning on HMI
GPU load close to maximum | High GPU load | Important computational load |  | GPU load Warning | 1. Follow up GPU load | Raise warning on HMI
GPU overload | Maximum GPU reached. Risk of latency increase | Very important computational load |  | GPU overload Warning | 1. Follow up GPU load | Raise warning on HMI
Frame lost | Lost frames detected during acquisition | Troubleshooting during image transfet between camera and DAQ Fast Controller |  | Acquisition Warning | 1. Image Post-processing to assess the impact | Raise warning on HMI.