## Measurement
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-ME-1 | Operator from HMI (Boy screens or appropriate tools) can interactively change plot parameters for all use cases of that particular diagnostic. |  |  | Yes
AC-ME-2 | Operator from HMI can crosscheck physics measurements from different diagnostics (including combined diagnostics measurement processing in D1 Control Group). |  |  | Yes
AC-ME-3 | Operator from HMI can crosscheck physics measurements from different diagnostics (including combined diagnostics measurement processing in D1 Control Group). |  |  | Yes

## Data Acquisition and Processing
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-DP-1 | Operator can verify that all processed data is correctly archived in ITER DB. |  |  | Yes
AC-DP-2 | Operator can read the Quality tag of a measurement from the HMI and interpret its details the metadata. |  |  | Yes
AC-DP-3 | Operator can verify the correct functioning of the signal chain from DAQ to CODAC networks. |  |  | Yes
AC-DP-4 | Operator can verify the correctness of DAQ data conversion by analysing the archived data and comparing it with a known signal. |  |  | Yes
AC-DP-5 | In case of problems with the data acquisition, Operator can correctly identify the location and the nature of the problem. |  |  | Yes
AC-DP-6 | Operator can verify that the time stamps of all DAQ and DP data on PON, SDN, and DAN are correctly reflecting the time of the observed plasma or other parameter. |  |  | Yes
AC-DP-7 | Operator can verify that the pre-processing algorithms are correctly executed in the DAQ device/controller. |  |  | Yes
AC-DP-8 | Operator can verify that the DAQ data is correctly transferred to the CPU with the required performance. Problems with the data transfer are reported and documented. |  |  | Yes
AC-DP-9 | As applicable, Operator can verify that the data processing is correctly synchronized with all data acquisition updates. This includes data processing from different sources and/or different update rates. |  |  | Yes
AC-DP-10 | In case of problems with the data processing chain, Operator can correctly identify the location and the nature of the problem in the data processing chain. |  |  | Yes
AC-DP-11 | Operator can verify the availability and correctness of the measurement errors related to physics measurements. |  |  | Yes
AC-DP-12 | Operator can verify all necessary metadata required to determine the production state, quality tag for the final measurement. |  |  | Yes
AC-DP-13 | Operator can verify that the algorithms for data processing are correctly executed. For this purpose all data along the processing chain is available on demand in ITER-DB for verification of the individual algorithms of the data processing chain |  |  | Yes
AC-DP-14 | Operator can monitor the latencies, jiiter and data throughput along the whole signal chain. |  |  | Yes

## Archiving
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-AR-1 | Operator can access all data (including raw data) acquired and stored in ITER DB during any operation including maintenance and |  |  | Yes
AC-AR-2 | Operator can analyse and validate the following on the acquired data: 1) time correlations between data streams 2) data integrity (values in range, saturations, timing, digits consistency, lower dynamics, rounding problems, etc) |  |  | Yes
AC-AR-3 | Operator can retrieve all data from FAT/SAT tests, for reproducing the test results and/or further analysis. This includes metadata relevant for the tests. |  |  | Yes

## Timing
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-TM-1 | Operator can verify that resetting/recovering DAQ during pulse must restore correct timing. |  |  | Yes
AC-TM-2 | Operator can verify that loss of TCN is detected and recovered through exception handling procedure. |  |  | Yes
AC-TM-3 | Operator can verify that the time stamps of archived data from PON, SDN and DAN are consistent within and between plant I&Cs. Correlation plots must have a timing accuracy of 50 ns (or as applicable). |  |  | Yes
AC-TM-4 | Operator can verify that the measurement time stamps reflect the time at which the measured plasma parameter was observed. |  |  | Yes
AC-TM-5 | Operator can monitor precision and accuracy of timing (clocks, triggers, time stamps). This information is also reflected in the meta data of quality tag. |  |  | Yes

## Configuration
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-CO-1 | Operator can manage Plant I&C configuration using Central I&C configuration tools. These include: 1) Prepare/Check/Load/Apply/Verify Configuration Objects 2) Check configuration integrity of Plant I&C by reading back current values and comparing to expected ones |  |  | Yes
AC-CO-2 | Operator can verify that after cutting power (class IV) to a cubicle the power-on procedures allow to establish exactly the same operating conditions (not limited to configuration objects) as it was just before the power cut. |  |  | Yes
AC-CO-3 | Operator can verify that a graceful shutdown after cutting class III power (with 5-10 minutes of UPS power available for the shutdown) is executed correctly. A bump-less restart (power-on) after power is re-established allows to establish the same operating conditions (not limited to configuration objects) as it was just before the graceful shutdown. |  |  | Yes
AC-CO-4 | Operator can access previous configurations by time or pulse number from ITER DB. |  |  | Yes
AC-CO-5 | Operator can access on the HMI, the status of condition flags which provide status of all configuration aspects of the various Plant I&C components. |  |  | Yes
AC-CO-6 | Operator can verify that manual interventions are correctly reflected in the system identfication status screens. |  |  | Yes
AC-CO-7 | Operator can use configuration management for executing all operation, testing, troubleshooting and maintenance procedures. Configuration management is accomplished through central CODAC tools including SUP-CONFIG, PSPS, and operational sequencer. Note 1: Testing procedures usually requires complex sequencing of frequent configuration changes. Note 2: Configuration management includes configuration transformation, verification and validation and handling of configuration errors while observing operating limits. |  |  | Yes
AC-CO-8 | Operator can monitor and assist configuration exception management procedures. These are the procedures which support detection and handling of configuration issues. Depending on the severity of the configuration errors operation may continue or may have to be terminated. |  |  | Yes
AC-CO-9 | Operator is able to verify that non-PV related configurations (e.g. firmware, non standard COTS software, cabling, faulty hardware, malfunctioning software,….) are appropriately recorded using ITER provided tools. |  |  | Yes
AC-CO-10 | Operator to verify the correct power up of the DAQ and COTS devices. Errors if any must be reported correctly. |  |  | Yes

## Alarm
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-AL-1 | Operator can monitor all summary alarms on HMI and navigate through the alarm tree to the lowest level alarm detail using CODAC alarm handler tool. |  |  | Yes
AC-AL-2 | Operator can verify the correctness of generated alarms by tracing signal chain to the source of the problem. Through this he can also verify that there are no false alarms generated. |  |  | Yes
AC-AL-3 | Operator can verify that all alarms provide guidance on the suggested corrective actions that the operator can take (if any). Operator is informed of the consequences of not treating the alarm properly. |  |  | Yes
AC-AL-4 | Operator can verify that the alarm handler is configured for efficient alarm handling. Note: Alarm trees must be designed to support hierarchical structure of alarms which allow to drill down from high level summary alarms to the detailed alarm input at the lowest level. |  |  | Yes
AC-AL-5 | Operator can verify that he can change the alarm thresholds to manage number of alarms generated. |  |  | Yes
AC-AL-6 | Operator can perform first fault analysis. This allows troubleshooting of the sequence in which a chain of alarms occurred. Accuracy of timestamps of alarms sufficient to support first fault identification. |  |  | Yes

## Calibration
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-CA-1 | Operator can perform all types of calibration for all calibration scenarios of the diagnostic. Note1 : Perform includes Prepare/Check/Load/Apply/Verify Note2: Types include online (measurements), offline (data analysis), in-situ (in dedicated test set up) and information from other sources. |  |  | Yes
AC-CA-2 | Operator verifies that he/she is able to manage calibration data in all repositories as required for all diagnostic use cases. |  |  | Yes

## Maintenance&Maintenability
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-MM-1 | Operator can verify, monitor and execute all self-tests and view test results on the HMI. |  |  | Yes
AC-MM-2 | Continuous availability studies during all acceptance phases are supported by plant I&C functions and procedures and the results of the availability studies are documented for the plant I&C (for example in the e-log book, test reports, and in the issue tracker) Operator can access these records and analysis results. Note: This includes the data from extended burn in (from weeks to months) before SAT. |  |  | Yes
AC-MM-3 | Operator can verify the effectiveness of risk mitigation measures as documented in RAMI report and implemented in the Plant I&C. |  |  | Yes
AC-MM-4 | Operator is able to introduce typical problems in the plant I&C (lose/disconnected cable, defect DAQ board, disconnected network (CODAC or PLN), a few wrong configuration parameters, etc.) and evalutate the troubleshooting capability of the system. Note: The plant I&C maintenance procedures and the related functions allow to determine all problems within a reasonable time (few minutes). The troubleshooting can be performed from the plant I&C HMI and using off-line analysis tools. |  |  | Yes
AC-MM-5 | Operator can verify that replacing a hardware component (component, device, subsystems) can be performed in a reasonable time (minutes to hours depending on complexity). Note: This includes power down/hot swap as applicable, finding replacement components in the inventory, requesting access to diagnostics building, power-on after replacement, re-testing, and documentation of the HW replacement. HMI (local or in control room) will be used to monitor the system status before, during and after the replacement. |  |  | Yes
AC-MM-6 | Operator can witness and verify the successful demonstration of software upgrade (adding a few selected additional functions) and software update (e.g. loading the firmware on a DAQ device or special software on a COTS controller. |  |  | Yes
AC-MM-7 | Operator can verify that maintenance activities (including those performed before acceptance testing) are documented (e.g. e-log book), the relevant data archived (as far as possible in ITER-DB, e-logbook and other repositories) and off-line analysis scripts for these maintenance activities are accessible and can be repeated to verify the results. |  |  | Yes
AC-MM-8 | Operator can access all maintenance related documentation. Note: This includes maintenance manual, maintenance plan with all relevant maintenance scenarios, previous maintenance reports (for e.g. in e-log book), issue tracker reports from previous bug fixes, RAMI analysis, failure statistics, etc. |  |  | Yes
AC-MM-9 | Operator can verify the existence of a minimum equipment list (MEL) and the documentation of operating limits. Note: The operations manual indicates clearly which failures and performance degradations are acceptable for continued operation and which failures and performance degradations require the termination of operation. |  |  | Yes
AC-MM-10 | Operator can operate (as applicable) the special maintenance equipment and tools delivered alongwith the Plant I&C. Note: This includes special hardware for maintenance and special software required the maintenance of the system like for example for firmware upgrades/updates. |  |  | Yes

## Operation
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-OP-1 | Operator can verify the operation of Plant I&C from Central I&C through COS:PSOS mapping. |  |  | Yes
AC-OP-2 | Operator can access and verify the successful logging of log messages and events, specifically the ones relevant for maintenance and troubleshooting. |  |  | Yes
AC-OP-3 | Operator can monitor all PON variables from the HMI for inspection and troubleshooting purposes. |  |  | Yes
AC-OP-4 | Operator can read from HMI the status of all the components of the system (through for example - alarms, status, health conditons, event) for full situational awareness of what is going on in the system. |  |  | Yes
AC-OP-5 | Operator can verify that the timing management executes correct timing sequence to ensure diagnostic’s participation in the pulse. |  |  | Yes
AC-OP-6 | Operator can verify correct implementation of exception detection and handling through appropriate procedures and functions. |  |  | Yes

## Availability
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-AV-1 | Operator can verify that the all documentation of plant I&C availability related topics is complete and approved. Note: This includes RAMI of I&C, MTBF and MTTR data for all plant I&C subsystems, controllers, devices and components as applicable. |  |  | Yes
AC-AV-2 | Operator can verify that the obsolescence management plan can be executed and is adequate for required availability. |  |  | Yes
AC-AV-3 | Operator can access the documentation for understanding of availability in presence of complete and partial failures, performance degradation, operator errors and configuration errors. |  |  | yes

## Performances
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-PE-1 | Operator is able to monitor actual performance together with performance limits and indications of exceeding these limits when changing parameters. Note: Example is the change of clock frequency and/or number of active channels leading to increase/decrease of data throughput. |  |  | Yes
AC-PE-2 | Operator can see on HMI the actual hardware configuration of the system and verify that the system has additional margins for slots and I/O channels as per PCDH requirements. |  |  | Yes
AC-PE-3 | Operator can verify that the system meets ITER electrical engineering requirements through inspection of documentation and/or the hardware. |  |  | Yes
AC-PE-4 | Operator can verify that all measurement requirements are met. Note: This is done through measurement performance monitoring and related quality tag and production state. |  |  | Yes
AC-PE-5 | Operator can monitor data throughput performance for the worst case scenario in which the maximum possible raw data stream is generated and archived. He/She can monitor performance parameters including data errors on communication links and validate that the archived data has been correctly and completely archived. |  |  | Yes
AC-PE-6 | Operator can assess the computational resources by monitoring the CPU/GPU load for each controller under worst case and typical operational case conditions. |  |  | Yes
AC-PE-7 | Operator can monitor on the HMI measurement noise (precision) and measurement error (accuracy) along the whole signal chain. He/She can compare measured noise and error with expected values. |  |  | Yes
AC-PE-8 | Operator can monitor the heat load in the cubicles. Note: This measurement can be an estimation based on corelation between heat load and power, temperature, fan status and other related parameters measured. |  |  | Yes

## Signal Conditioning
Criteria Name and ID
ID | Acceptance Criteria | Notes | Name | Applicability
AC-SC-1 | Operator will be able to check that the signal conditioning is sufficiently instrumented to monitor its health and performance. Note 1: It must include functions required for all relevant use cases (operation, calibration testing, maintenance, configuration) and corresponding detailed procedures. Note 2: Signal conditioning instrumentation variables must cover all relevant attributes including control, monitoring, configuration, health, conditions, warning, alarms, events, etc. |  |  | Yes
AC-SC-2 | Any failure or performance degradation of signal conditioning is reliably detected and reported to the operator through HMI. During acceptance testing operator can introduce failures or performance degradation through manual interventions (like disconnecting a cable or adding an attenuator) and/or using I&C functions (like controlling the voltage of a power supply or the signal levels on a function generator used for testing). |  |  | Yes

## Categories
ReqID | Measurement
| Req Name | Description | Source | Section
DX-XX-MER-XXX-XXX |  | Specific per PBS from SRD-55
DX-XX-DAQ-ADC-003 | Acquired data to Central I&C System | The plant system shall be able to send acquired information to the Central I&C system in raw data | 27LH2V | R97
DX-XX-DAQ-DPS-001 | Processed signals of the derived plasma | Computer hardware and software associated with a diagnostic sub-system shall provide: processed signals of the derived plasma | 28B39L | 55s863
DX-XX-DAQ-DPS-002 | I&C Data available for Logging | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for logging. | 27LH2V | R3
DX-XX-DAQ-DPS-003 | I&C Data available for Monitoring | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for monitoring. | 27LH2V | R3
ReqID | Data Acquisition and processing
| Req Name | Description | Source | Section
DX-XX-DAQ-ADC-001 | Acquire raw data | Plant system shall acquire raw data | 27LH2V | G12
DX-XX-DAQ-ADC-002 | Raw Measurement | Computer hardware and software associated with a diagnostic sub-system shall allow the recording of raw measurements of that diagnostic sub-system. | 28B39L | 55s862
DX-XX-DAQ-ADC-003 | Acquired data to Central I&C System | The plant system shall be able to send acquired information to the Central I&C system in raw data | 27LH2V | R97
DX-XX-DAQ-DPS-001 | Processed signals of the derived plasma | Computer hardware and software associated with a diagnostic sub-system shall provide: processed signals of the derived plasma | 28B39L | 55s863
DX-XX-DAQ-DPS-002 | I&C Data available for Logging | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for logging. | 27LH2V | R3
DX-XX-DAQ-DPS-003 | I&C Data available for Monitoring | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for monitoring. | 27LH2V | R3
DX-XX-DAQ-DPS-004 | Process monitoring function | The plant system I&C shall implement Process monitoring function. | 27LH2V | R70
DX-XX-DAQ-DPS-005 | Generation of data streams function | The plant system I&C shall implement Generation of data streams function. | 27LH2V | R70
DX-XX-DAQ-DPS-006 | I&C Data available for Analysis | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for analysis. | 27LH2V | R3
DX-XX-DAQ-DPS-007 | Processing Information | All information issued from the process shall be supplied with an identifier, a time stamp and a quality flag including error identification in case of error. Units and full name of the information may not be required in the dynamic data if defined in the associated static metadata. | 27LH2V | R71
DX-XX-DAQ-DPS-008 | Computed data to Central I&C System | The plant system shall be able to send computed information to the Central I&C system in engineering units. | 27LH2V | R97
DX-XX-DAQ-DPS-009 | Experimental data processing function | The plant system I&C shall implement Experimental data processing function | 27LH2V | R70
ReqID | Archiving
| Req Name | Description | Source | Section
DX-XX-DAQ-ARC-001 | I&C Data available for Archiving | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for archiving. | 27LH2V | R3
ReqID | Timing
| Req Name | Description | Source | Section
DX-XX-DAQ-TIM-001 | Synchronized with ITER central time reference | The plant system I&C shall be synchronized with ITER central time reference. | 27LH2V | R90
ReqID | Configuration
| Req Name | Description | Source | Section
DX-XX-CFG-PCG-001 | Configuration function | The plant system I&C shall implement Configuration function | 27LH2V | R70
ReqID | Alarms
| Req Name | Description | Source | Section
DX-XX-ALR-ERR-001 | Error and trace logging function | The plant system I&C shall implement Error and trace logging function | 27LH2V | R70
DX-XX-ALR-OAL-001 | Diagnostics Alarm | Each diagnostic shall provide signals for monitoring its operational state or conditions affecting other systems consistent with the operation of the CODAC system | 28B39L | 55s125
DX-XX-ALR-OAL-002 | Alarms function | The plant system I&C shall implement Alarms function | 27LH2V | R70
DX-XX-ALR-OAL-003 | Maintain active alarm status | The plant system I&C shall maintain the status of all active alarms and shall transmit any change of this status (alarm raised, alarm cleared). | 27LH2V | R81
DX-XX-ALR-OAL-004 | Status Information of Alarm Conditions | Plant system I&C shall provide status information for alarm conditions that is required to operate the plant system I&C from Main Control Room (MCR). | 27LH2V | R4
DX-XX-ALR-OAL-005 | Status Information of Trip Conditions and corrective actions | Plant system I&C shall provide status information for trip conditions and corrective actions that is required to operate the plant system I&C from Main Control Room (MCR). | 27LH2V | R4
DX-XX-ALR-OAL-006 | Enable Alarm Reduction | The alarm shall carry information to the CODAC system to enable alarm reduction (not applicable to PSS). | 27LH2V | R82
DX-XX-ALR-ERR-001 | Error and trace logging function | The plant system I&C shall implement Error and trace logging function | 27LH2V | R70
ReqID | Calibration
| Req Name | Description | Source | Section
DX-XX-CAL-PCL-001 | Calibration Function | Computer hardware and software associated with diagnostic sub-systems shall be able to provide calibration functions | 28B39L | 55s864
DX-XX-CAL-PCL-002 | Calibration factor | Calibration factor and conversion formula shall be configurable. | 27LH2V | R73
ReqID | Operation
| Req Name | Description | Source | Section
DX-XX-OPR-AUT-001 | Management of events function | The plant system I&C shall implement Management of events function. | 27LH2V | R70
DX-XX-OPR-AUT-002 | Implement COS | Plant system I&C shall implement COS. | 27LH2V | R122
DX-XX-OPR-AUT-003 | Implement PSOS | Plant system I&C shall implement PSOS. | 27LH2V | R122
DX-XX-OPR-AUT-004 | Status Information of COS | Plant system I&C shall provide status information for Common Operating States (COS) that is required to operate the plant system I&C from Main Control Room (MCR). | 27LH2V | R4
DX-XX-OPR-AUT-005 | Status Information of PSOS | Plant system I&C shall provide status information for Plant System Operating States (PSOS) that is required to operate the plant system I&C from Main Control Room (MCR). | 27LH2V | R4
DX-XX-OPR-EXP-001 | Built in Time-out function | Time critical devices shall have built-in time-outs to ensure correct operation in case of Central I&C Systems failure. | 27LH2V | R15
DX-XX-OPR-EOP-001 | Log Message information | Plant I&C shall implement Logg Message function.A log message shall include: - A time stamp; - A process identifier according to the naming scheme; - A text explaining the event; - A message level (debug, info, warning, error). | 27LH2V | R85
DX-XX-OPR-EOP-002 | Log Message Qualifiers | The following log messages shall be recorded with their qualifiers in the logging system: - All timing, PSH, plant system Controller, PLC or embedded system events or state changes; - All operations related to data configuration (creation/modification/deletions of variables, threshold change); - All transitions in operating states; - All commands sent by Central I&C systems; - All binary state changes (e.g. valve opened or closed); - All events concerning an analogue variable or a group of analogue variables (threshold overshooting, out of range, discrepancy); - All variable validity changes; - All actions done locally by operators (log on/off, local commands, variable tagging or forcing); - All local alarm acknowledgements. | 27LH2V | R86
DX-XX-OPR-EOP-003 | Event Reporting | The plant system I&C events shall be reported in the logging and also alarms. This information shall also be propagated to the Central I&C system. | 27LH2V | R94
DX-XX-OPR-HMI-001 | I&C Data available for Visualization | Plant system I&C shall make available all data acquired from sensors, with a time stamp, to Central I&C Systems for visualization. | 27LH2V | R3
DX-XX-OPR-TOK-001 | control tokamak system operations | The Diagnostic system shall contain instrumentation to ensure the following functions: To control tokamak system operations | 28B39L | 55s1801
DX-XX-OPR-TOK-002 | prepare plasma pulses | The Diagnostic system shall contain instrumentation to ensure the following functions: prepare plasma pulses | 28B39L | 55s1801
DX-XX-OPR-TOK-003 | ensure plasma discharge quality | The Diagnostic system shall contain instrumentation to ensure the following functions:To ensure plasma discharge quality | 28B39L | 55s1801
ReqID | Maintenance and Maintainability
| Req Name | Description | Source | Section
DX-XX-MTN-HLT-001 | Include self-tests and live tests | Plant system monitoring shall include self-tests and live tests. | 27LH2V | R96
DX-XX-MTN-HLT-002 | Equipment self diagnostics | Computers and equipment shall have provisions for self-diagnosis and provide a visual indication of the status on the local front panels and at the MCR | 27LH2V | G10
DX-XX-MTN-HLT-003 | Equipment self-check | Computers and equipment should repeat self-checks at scheduled times. | 27LH2V | G10
DX-XX-MTN-MAN-001 | Maintenance procedures and tools | Maintenance procedures and tools for diagnostics shall be provided for maintenance scenarios with a probability exceeding 10^-1 over the 20 year life of ITER | 28B39L | 55s2777
DX-XX-MTN-SMG-001 | System management function | The plant system I&C shall implement System management function | 27LH2V | R70
DX-XX-MTN-SMG-002 | Remote control functions | Remote control functions shall be available (reboot, configure, start, stop, switch to local / central control mode). These functions shall comply with the safety rules of the ITER site. | 27LH2V | R87
DX-XX-MTN-SMG-003 | Status Information of power supply status | Plant system I&C shall provide status information for power supply status that is required to operate the plant system I&C from Main Control Room (MCR). | 27LH2V | R4
DX-XX-MTN-SMG-004 | Cubicle Monitoring | The I&C cubicles shall be equipped with a monitoring system for doors, temperature and cooling monitoring | 27LH2V | R157
DX-XX-MTN-SMG-005 | Plant I&C functions | The plant I&C systems shall provide those functions that are necessary to operate, to control and to monitor the underlying plant systems, to provide investment protection and to protect people and the environment | 28C2HL | 45s73
DX-XX-MTN-SMG-006 | Monitor Equipment Performance | Information on equipment performance shall be monitored. Performance information such as field bus, CPU load and memory usage or network bandwidth utilisation shall be recorded. | 27LH2V | R93
DX-XX-NMT-ANL-003 | Component naming | A convention for uniquely identifying parts and components for ITER is defined in the ITER Numbering System for Parts/Components, see [RD3]. This naming convention is applicable to any component of the plant system I&C. | 27LH2V | R65
DX-XX-NMT-ANL-004 | Maintenance Requirement | Maintenance requirements and operations shall conform to the ITER scheduled maintenance periods and scenarios. | 28B39L | 55s972
DX-XX-NMT-ANL-005 | Minimize Maintenance | Diagnostic Sub-systems shall be designed where possible to be easily maintained | 28B39L | 55s283
DX-XX-NMT-ANL-006 | Maintenance Plan | Maintenance plans shall be developed for all maintenance scenarios with a probability of occurrence greater than 10-6, over the 20-year life of ITER. The maintenance plans shall be developed during the design phase and submitted to IO for approval at each system review. | 28B39L | 55s2776
DX-XX-NMT-ANL-007 | Time to detect, identify, locate and repair failures | Diagnostics systems shall be designed to reduce the time to detect, identify, locate and repair failures, or sign of impending failure | 28B39L | 55s2399
DX-XX-NMT-ANL-008 | Accessibility to the physical interface points | Accessibility to the physical interface points shall be taken into account as a way to improve the maintainability and inspectability of the concerned components, and thus ultimately the availability of the systems. | 28B39L | 55s2400
DX-XX-NMT-ANL-009 | Diagnostic sub-systems and components identified on the RAMI analyses | Diagnostic sub-systems and components identified on the RAMI analyses as requiring more frequent maintenance to fulfil the required system availability shall be integrated to allow easy maintenance aiming at short term Maintenance. | 28B39L | 55s2404
DX-XX-NMT-ANL-010 | RAMI outputs | RAMI outputs shall be taken into account during the design, testing, operation and maintenance phases of the Diagnostics Systems. | 28B39L | 55s2383
DX-XX-NMT-ANL-011 | Diagnostics systems RAMI analysis | Diagnostics systems RAMI analysis shall be performed for each sub-system prior to the Preliminary Design Review. | 28B39L | 55s1434
DX-XX-NMT-CHG-001 | Obsolescence Management | It shall be possible to replace plant system I&C equipment to cope with I&C maintenance issues, plant system I&C upgrades, I&C hardware or software obsolescence, or as a result of it becoming increasingly expensive to operate and maintain | 27LH2V | 3.4.9
DX-XX-NMT-CHG-002 | Obsolescence Management Plan | A proactive management plan for obsolescence describing the strategies for identification and mitigation of the effects of obsolescence throughout all stages of I&C life cycle. This management plan shall be produced during the design phase and maintained through all the phases. | 27LH2V | D71
DX-XX-NMT-CHG-003 | Obsolescence Management - New Equipment | Every new I&C equipment shall be documented in the same way as was required for the initial procurement. | 27LH2V | R38
DX-XX-NMT-CHG-004 | Spare parts for non-standard equipment | The plant system responsible officer (and plant system I&C supplier, if appropriate) has to provide and pay for special integration and additional maintenance including spare parts for non-standard equipment. | 27LH2V | R283
DX-XX-NMT-CHG-005 | Manage I&C | CODAC shall manage instrumentation and control (I&C) maintenance, plant system upgrades, and I&C hardware and software technology evolution. | 28C2HL | 45s664
DX-XX-NMT-CHG-007 | Hot Swapping | Hot swapping shall be used whenever it is required by the RAMI analysis of the plant system | 27LH2V | G8
DX-XX-NMT-CHG-008 | Obsolescence Management | CODAC shall provide obsolesce management, using versioning control for hardware and software. | 28C2HL | 45s665
DX-XX-NMT-CHG-010 | Minimum time to replace | The maintainability shall ensure the minimum time to recognize, isolate and correct a malfunction, to understand and apply technical procedures, to gain access to faulty items, to repair or replace faulty items and to test and verify accuracy and adequacy of the maintenance actions | 28B39L | 55s2401
DX-XX-NMT-CHG-011 | Interchangeable components | Diagnostic system components that need to be replaced throughout the lifetime of the ITER campaign shall be designed or chosen to be interchangeable. | 28B39L | 55s2393
DX-XX-NMT-STB-002 | Adequate working/maintenance policies | Adequate working/maintenance policies shall be demonstrated to ensure that hazardous products will not affect workers during normal operations and maintenance procedures. | 28B39L | 55s2428
DX-XX-NMT-TES-001 | I&C computing and network resources information | The acceptance tests shall provide information about the load factor of the plant I&C computing and network resources. | 27LH2V | R104
ReqID | Availability
| Req Name | Description | Source | Section
DX-XX-NAV-AVA-001 | Obsolescence Management | The supplier shall inform ITER Organization of any risk regarding components’ obsolescence, and make all necessary recommendation to fix that risk | 28B39L | 55s1024
DX-XX-NAV-AVA-002 | Measurement Parameter Availability - Machine Protection | Measurement parameters for Machine Protection (Investment Protection (IP)) shall have redundancy via a set of diagnostic systems to ensure 98.5 % Availability is achieved. | 28B39L | 55s1438
DX-XX-NAV-AVA-003 | Measurement Parameter Availability - Basic Control | Measurement parameters for Basic Control shall be have redundancy via a set of diagnostic systems to ensure 98% Availability is achieved. | 28B39L | 55s1440
DX-XX-NAV-AVA-004 | Measurement Parameter Availability - Advanced Control | Measurement parameters for Advanced Control shall have redundancy via a set of diagnostic systems to ensure 91% Availability is achieved in the Hydrogen phase and 90 % in the DT phase | 28B39L | 55s666
DX-XX-NAV-AVA-005 | Measurement Parameter Availability - Physics | Measurement parameters for performance and physics shall have redundancy via a set of diagnostic systems to ensure 90% Availability is achieved in the Hydrogen phase and 89% in the DT phase. | 28B39L | 55s2531
DX-XX-NAV-AVA-006 | Availability at D-T Phase | Diagnostics systems shall be designed for an active (D-T) phase lasting at least 14 years. | 28B39L | 55s2318
DX-XX-NAV-AVA-007 | Parameter to measure availability | Diagnostic systems shall contain instrumentation to measure all parameters that may affect their system, availability or which may predict failures or indicate the need of maintenance. | 28B39L | 55s1802
DX-XX-NAV-AVA-008 | Minimise the time to repair | Diagnostic systems shall be designed to minimise the time to repair providing optimum Availability. | 28B39L | 55s2784
DX-XX-NAV-AVA-009 | Availability Compliance with RAMI | The availability of the plant systems I&C shall be compliant with the RAMI requirements of the plant system. | 27LH2V | R101
DX-XX-NAV-AVA-010 | Operation Capabilities - 30000 Pulses | Diagnostics systems shall be capable of operating for at least 30,000 pulses in order to satisfy the overall requirement specified in the Project Requirement. | 28B39L | 55s2315
DX-XX-NAV-AVA-011 | Operating for 11 Days | Diagnostics systems shall be designed to be capable of operating for periods of 11 consecutive days while accommodating three-shift daily plasma operation, followed by 3 days of routine maintenance | 28B39L | 55s2319
DX-XX-NAV-AVA-012 | 16 Month Continuous Operation | Diagnostics systems shall be designed so that plasma operation can be conducted for periods of up to 16 months continuously in three 8h work-shift daily operating mode to perform the following actions: plasma operations, test, conditioning, routine maintenance. | 28B39L | 55s2320
DX-XX-NAV-REC-001 | Diagnose faults and facilitate fast recovery | The plant system I&C shall be monitored in a homogeneous way in order to diagnose faults and facilitate fast recovery. | 27LH2V | R88
DX-XX-NAV-REC-002 | Software Version Control | The software versioning control tool shall be Subversion. | 27LH2V | R115
DX-XX-NAV-REC-003 | Version Control for Documents | All documentation shall be under version control | 27LH2V | R45
DX-XX-NAV-REC-004 | Deliverables Backup | All deliverables in electronic format shall be backed up after the acceptance phase in order to secure a functional restore state. | 27LH2V | R21
DX-XX-NAV-REL-001 | Redundancy | Redundancy shall be used whenever it is required by the RAMI analysis of the plant system | 27LH2V | G9
DX-XX-NAV-REL-002 | Online failure rate | Diagnostic sets providing measurement parameters for Machine Protection (1a1) will be designed for a combined online failure rate of less than 5*10^-5 per measurement per operational (calendar) year. | 28B39L | 55s1439
DX-XX-NAV-REL-003 | Machine Protection Redundancy of at least 2 | Diagnostic sets providing measurement parameters for Machine Protection (1a1) shall have redundancy of at least 2 (for method, type or location), and >120% for component count where appropriate | 28B39L | 55s664
DX-XX-NAV-REL-004 | Online failure rate of less than 5E-3 | Diagnostic sets providing measurement parameters for Basic Control (1a2) will be designed for a combined online failure rate of less than 5*10^-3 per measurement per operational (calendar) year. | 28B39L | 55s1441
DX-XX-NAV-REL-005 | Basic Control Redundancy of at least 2 | Diagnostic sets providing measurement parameters for Basic Control (1a2) shall have redundancy of at least 2 (for method, type or location) and >100% for component count, where appropriate | 28B39L | 55s665
DX-XX-NAV-REL-006 | Online failure rate of less than 0.1 | Diagnostic sets providing measurement parameters for Advanced Control (1b) shall be designed for a combined online failure rate of less than 0.1 per measurement per operational (calendar) year. | 28B39L | 55s666
DX-XX-NAV-REL-007 | Measurement Parameter evaluation | Diagnostic sets providing measurement parameters for Evaluation & Physics Studies (2) will be designed for a combined online failure rate of less than 0.1 per operational year. | 28B39L | 55s668
ReqID | Performance
| Req Name | Description | Source | Section
DX-XX-NPR-EFF-001 | Alarm Philosophy - Performance | The core principles underline this alarm philosophy - Performance monitoring: the performance of the alarm system should be assessed during design and commissioning to ensure that it is usable and effective under all operating conditions. Regular auditing should be continued throughout the plant system life to confirm that good performance is maintained, | 27LH2V | R361
DX-XX-NPR-SCE-001 | Sufficient margin to add functions | The plant I&C computing and network resources shall provide sufficient margins to allow to add additional functions and increased data rates. | 27LH2V | R103
DX-XX-NPR-SCE-002 | Reserve Hardware slot | Additional reserve slots (not equipped) per backplane type shall be more than 20%. | 27LH2V | R105
DX-XX-NPR-SCE-003 | Reserve Not equipped IO channels | Additional reserve I/O channels (not equipped) per type shall be more than 20%. | 27LH2V | R106
DX-XX-NPR-SCE-004 | Reserve equipped IO channels | Additional reserve I/O channels (equipped) per type shall be more than 5%. | 27LH2V | R107
DX-XX-NPR-SCE-005 | PCIe for IO expansion | PCI Express shall be used in the I&C System for I/O expansion. | 333K4C | G1-20
DX-XX-NPR-SCL-001 | Earth / ground | Neither the MESH-CBN, nor any part of it, or anything described as an “earth” or “ground” shall ever be used as an electrical return path for any signals, control or instrumentation of any type; or as a return path for DC or AC power at any voltage or current level. | 4B523E | 8.2
DX-XX-NPR-SCL-002 | AC power distribution | AC power distribution - Incoming power cable /cables from distribution or sub-distribution boards shall have separate terminals inside the cubicle. | 4H5DW6 | 5.6
DX-XX-NPR-SCL-003 | Be protected with adequate current protection | Each distribution circuit must be protected with adequate current protection | 4H5DW6 | 5.6
DX-XX-NPR-SCL-004 | Have inbuilt operating and overload mechanisms | The circuit breakers must have a tripping curve appropriate for the features of the protected circuit. Circuit breaker shall be double pole and shall have inbuilt operating and overload mechanisms. | 4H5DW6 | 5.6
DX-XX-NPR-SCL-005 | Power socket for fast controllers | Power socket/strips for fast controllers/PSH shall be protected by circuit breaker of appropriate rating. | 4H5DW6 | 5.6
DX-XX-NPR-SPL-002 | Sensor Information update rate | Duration for update of information from sensors to the Plant Operation Network shall be less than 1 sec (for PSS, this is only applicable to communication between PSS and CSS). | 27LH2V | R108
DX-XX-NPR-SPL-003 | Command to actuators | Duration for unitary commands from Central I&C networks to actuators shall not exceed 200 ms. | 27LH2V | R109
DX-XX-NPR-SPL-004 | Diagnostics performance requirement | Plant system I&C participating in the diagnostics or plasma feedback control shall have specific performance requirements (not applicable to PSS). | 27LH2V | R110
DX-XX-NPR-PRA-001 | Diagnostics Parameters -1 | The Range or coverage, resolution, accuracy parameters applicable to diagnostics shall be defined during the design phase | 28B39L | 55s1129
DX-XX-NPR-PRA-002 | Diagnostics Parameters -2 | The Range or coverage, resolution, accuracy parameters applicable to diagnostics shall be Approved before FDR | 28B39L | 55s1129
DX-XX-NPR-PRA-003 | Diagnostics Parameters -3 | The Range or coverage, resolution, accuracy parameters applicable to diagnostics shall be defined based on their role, contribution and plasma conditions | 28B39L | 55s1129
ReqID | Signal Conditioning
| Req Name | Description | Source | Section
Covered through various requirements | Signal Conditioning | Signal Conditioning must be adequately instrumented to support all other relevant requirements categories. | 27LH2V | various
MAIN TOPICS TO COVER (CONFIGURATION)
Configuration Objects
Transformation and verification/validation algorithms
Sequence for maintenance, troubleshooting, testing
Procedures and HMI
Exception detection and handling
Configuration history
Non PV configuration ( firmware, non standard COTS software, cabling, HW/SW redundancy usage, ….)
MAIN TOPICS TO COVER (ALARMS)
Multiple alarm/warning trees
Severity and priority (warnings, minor and major alarms)
Managing alarm thresholds and masking alarms
Acknowledging alarms
Managing alarm rates (reduce rate)
Alarm actions
MAIN TOPICS TO COVER (Maintenance/Maintainability)
Troubleshoot (signal chain, subsystems, device, components)
Repair HW (components, devices, subsystems, cabling, etc.))
Replace HW components
Update/upgrade software
Re-testing / re-commissioning
Scheduling of maintenance
Documentation of maintenance activities
Inventory management
On-line / Off-line analysis for maintenance
Health monitoring
Self-testing
Performing availability studies
Availability study results, failure statistics, RAMI analysis
Other : In-situ, off-site, preventative Failure detection, performance degradation detection
Minimum equipment list ( lists which equipment can be failed but operation is possible)
Special maintenance equipment (not part of plant I&C)
Central CODAC integration
Interplant operation
Remote (expert) participation
Maintenance procedures
Maintenance manual
Operating limitations
Equipment protection
Maintenance modes, states and use cases defined and documented.
Development system available in XPOZ as required for maintenance
MAIN TOPICS TO COVER (Availability)
RAMI Analysis, MTBF and MTTR for subsystems, devices and components
Availability studies, measurement of availability
Environmental conditions
Exception detection and handling
Definition of availability for each diagnostics and plant I&C
Understanding of availability and the operating limitations in presence of partial failures and/or performance degradations
Obsolescence management
MAIN TOPICS TO COVER (Performances)
Data throughput
Raw data archiving
Latency
Computatinal capability
Redundancy
Accuracy and precision
Data error on communication links
Heat load in cubicle
Signal chain integrity
MAIN TOPICS TO COVER (Signal Conditioning)
Instrumentation of signal conditioning with control, monitor, configuration, condition, alarm, calibration, and system management functions
Failure and performance degradation reporting
HMI for signal conditioning
MAIN TOPICS TO COVER (DAQ and Processing)
Device support incl. DAQ system and related device state-machine
ADC and DAC timing signals
DAQ on COTS devices/controllers
Pre-processing in FPGA
Algorithms
main processing in slow/fast controller CPU (input to final measurement algorithm)
data processing on COTS controllers
Data coming from different sources
Algorithms
RTF
Off-line analysis for physics and engineering